[image: image1.jpg]—ULEX

European Union Rule of Law Mission

Kosovo





European Union Rule of Law Mission in Kosovo EULEX Kosovo

EULEX Kosovo

Ndertesa Farmed 

“Muharrem Fejza” p.n.

Lagja e Spitalit

10000 Pristina, Kosovo

Tel: +(381) 38 222 010 2000

www.eulex-kosovo.eu
CLARIFICATIONS (2)
PUBLICATION REFERENCE: PROC/117/09/Veterinary Equipment (EuropeAid/128823/C/SUP/KOS)
1. Questions regarding the “Technical Specifications” annex II + III.  

	Question

No
	Lot No
	Question
	Answer

	1
	Lot 4: Surgical Instruments & Autoclave Equipment
	Item No. 29 Autoclave Equipment:

Do you accept also Standard prEN 13060-4. This autoclave adheres to the requirements stipulated in the denomination "Class S". 


	The autoclave must be a “B type”. 

“S” type autoclaves are NOT compliant.

	2
	Lot 4: Surgical Instruments & Autoclave Equipment


	Item No. 29 Autoclave Equipment:

Is tap water connection necessary for autoclave equipment?


	No.

	3
	Lot 4: Surgical Instruments & Autoclave Equipment


	Item No. 30 Autoclave Nylon Film: 

Do you accept also autoclave bag 100 x 300 mm to be supplied in pieces?


	Yes.

	4
	Lot 4: Surgical Instruments & Autoclave Equipment


	Item No. 31 Autoclave Indicator Tape

Do you accept also autoclave tape 19 width?


	Yes, (requested dimension is approximate).

	5
	Lot 5:  Miscellaneous Veterinary Equipment


	Item No. 6 Clippers: 

Do you accept also clippers with two speeds?
	Yes.

	6

	Lot 5:  Miscellaneous Veterinary Equipment


	Item No. 15 Pre room/treatment/XRAY table: 

Do you accept also dimensions 560 x 1130 mm and feet 870 mm?


	Yes. The dimensions required in the specs are approximate and these are close enough. 

	7

	Lot 5:  Miscellaneous Veterinary Equipment
	Item No. 17 Surgical operating light: 

Is High intensity illumination - 2600-6900 lux/0.8 m enough?

	NO. In accordance with the technical specifications, are required approximately 50,000 light intensity at 1m distance. 

Products with a lower lux level will be declared NON compliant. 



	8
	Lot 5:  Miscellaneous Veterinary Equipment
	Item No. 20 Drapes:

Do you accept also dimensions 60 x 90 cm?


	Drapes required dimension is    approximately      75 x 90 cm plain, (e,g 60 X 90     is accepted).  

              








PAGE  
2

